## Slide 1
Evaluation of the compressive strength of bone cements with antibiotics
Authors: S. Medina-Garcia*a,b, O. A. González-Estradab, N. Molinaa, D. Rincón-Ovallea, A. Sáenz-Suarezb, I. Sáenz-Suarezb, R. E. Castañedaa, L. Ramíreza and L. E. Garzóna.
aLíneas Hospitalarias SAS
bUniversidad Industrial de Santander

## Slide 3
Fig. 1. surgical procedure for the placement of an intramedullary nail in the tibia[1].
Fig. 2. Uses of bone cement in medical treatments[2].

## Slide 4
Objectives
Determine how the compression strength of low-viscosity bone cement changes when mixed with different dosages and types of antibiotics, according to the ASTM F451 standard through compression tests, in order to establish the quantities of antibiotics that meet established regulations.

## Slide 5
Objectives
Identify the main antibiotics and dosages that are mixed with low-viscosity G3 bone cement, brand G21, in surgical practice with humans through a review of the state of the art and clinical staff.
Perform compression tests for the identified mixtures, following the ASTM F451 standard to obtain compression strength.
Analyze the mechanical property of compression strength for the bone cement and antibiotic mixtures, evaluating its variation based on the different applied antibiotic dosages.

## Slide 6
Methodology

## Slide 7
Experimental Method
Observation
Analysis of variables
Selection of data (quantitative data)
Control
Conclusion

## Slide 8
Materials

## Slide 9
Fig. 3. a) Vancomycin powder, b) Liquid gentamicin, and c) Low-viscosity G3 bone cement (brand G21). Source: author

## Slide 10
test tube size | cylinder | 
radius (r) | 3 | mm
length (l) | 12 | mm
volume(Vp) | 0,339 | ml
Total number of test tubes without antibiotic |  |  | 
test tube mass(Mp)= |  | Vp*ρ= | 0,475
#test tubes=total mass/Mp= |  | 111,96 | 
Table. 1. mass and volumen of bone cement G3 low viscosity. Source: author
Bone cement G3 Low viscosity |  | 
powder pmma | grams | milliliters
mass = | 34,4 | 
volume= |  | 17,99
Liquid monomer |  | 
mass-liquid= | 18,78 | 
volume-liquid= |  | 20
total mass= | 53,18 | 
total volume = |  | 37,99
Table. 2. Dimensions of the test tubes. source: author
Table 3. calculation of the total number of test tubes per envelope of G3 low viscosity bone cement. Source: author

## Slide 11
vancomycin 500mg | dose (grams) | representative sample (g) | amount of pmma (grams) | liquid volume (ml) | #test tubes
1 mix | 2 | 0,2 | 3,44 | 2 | 10
2 mix | 4 | 0,4 | 3,44 | 2 | 10
3 mix | 6 | 0,6 | 3,44 | 2 | 10
4 mix | 8 | 0,8 | 3,44 | 2 | 10
Table. 4. quantities of bone cement and antibiotic to make the test tubes. Source: author

## Slide 12
Equipment

## Slide 13
Fig. 4. Analytical balance with 0.6 grams of vancomycin powder. Source: author

## Slide 14
Fig. 5. a) UHMW mold for fabricating the test specimens, b) flat aluminum disk for mold adjustment, c) mold implementation, d) syringes as containers for mixing with the 316L steel pin. Source: author

## Slide 15
Fig. 6. Hydraulic press MTS Bionix. Source: author

## Slide 16
Results

## Slide 17
Fig. 7. Typical stress vs. strain curve in compression tests.

## Slide 18
Chart: Esfuerzo - deformación para cemento óseo con vancomicina
Graf. 1. Stress vs. strain curves for bone cement mixtures with vancomycin. Source: author

## Slide 19
Chart: Esfuerzo - deformación para cemento óseo con gentamicina
Graf. 2. Stress vs. strain curves for bone cement mixtures with gentamicin. Source: author

## Slide 20
Graf. 3. Compression strength of all samples. Source: author
Chart: Resistance to compression
| Without antibiotics | 2 grams | 4 grams | 6 grams | 8 grams | 1 ampoules | 2 ampoules | 3 ampoules | 6.25 ampoules
 | 92.24 | 85.464 | 64.646 | 78.5 | 72.986 | 63.6 | 58.44 | 53.83 | 33.496

## Slide 21
Analysis

## Slide 22
Conclusions
*None of the samples with gentamicin exceeded 70 MPa in compressive strength, indicating that the use of this antibiotic individually is not feasible, as it does not meet mechanical standards. *Regarding the samples with vancomycin, three of them exceeded 70 MPa, ranging between 85.46 MPa and 72.99 MPa for the samples with 2, 6, and 8 grams. The sample with 4 grams had a compressive strength of 64.65 MPa, suggesting a need to review or repeat this test due to a possible contact surface error. The other samples are viable for intraoperative use, as their compressive strength complies with the established standards. *The results of these tests show that using any liquid antibiotic significantly changes the mechanical properties of the bone cement, making it unfeasible for medical treatments as its mechanical integrity cannot be assured. *We recommend continuing this research by combining these antibiotics and retesting their compressive strength. Additionally, a study on the elution of the antibiotic in each sample is essential, as it is crucial to assess the antimicrobial behavior of this material besides its mechanical capacity.

## Slide 23
Bibliography
[1]. Hospital general Jose Maria Velazco Ibarra. (2021, 6 de octubre). #Napo | Se realiza con éxito cirugía traumatológica para la colocación de clavo endomedular en la tibia a paciente que sufrió accidente laboral. [Imagen]. Facebook. https://www.facebook.com/HJMVITena/posts/2842325389412710/?_rdr
[2]. Morejón, Lizette, Mendizábal, Eduardo, Delgado, José Ángel, Ledea, Oscar Ernesto, Fuentes, Gastón, López, Mónica, & Durán, Ivette. (2021). Cementos óseos acrílicos: desarrollo de nuevas formulaciones bioactivas y con actividad terapéutica. Anales de la Academia de Ciencias de Cuba, 11(1), . Epub 30 de agosto de 2021. Recuperado en 06 de agosto de 2024, de http://scielo.sld.cu/scielo.php?script=sci_arttext&pid=S2304-01062021000100010&lng=es&tlng=es.

## Slide 24
Thanks